GUIDELINES FOR PISCES OPERATION

(January 1, 2004)

POWER-UP SEQUENCE:

1. In the control room:

a.	Turn on:	PC

b.	Hit "F8” key after the SCSI lines have been printed, about the time that "DMI

Pool” is printed

c.	Hit #8 key to "boot previous version of DOS”

d.	Go to d: and then cd to \don

e. 	If desired, set date/time on PC

2. In the control room:

a.	Turn on:	Leach interface box

b.	Press:		"reset” button twice on Leach interface box

3. At telescope:

a.	Turn on:	Leach controller & wait 30 seconds

4. The F/O (fiber optic) light on Leach interface box should shine green

[The Power light doesn’t work]

5. On the PC, from the d:/don directory:

a.	Type "wfirc” to boot PISCES software

PRECAUTION in HIGH HUMIDITY:

1. At the telescope, place a “trouble light” as close as possible to the front of the dewar window to warm the metal and prevent condensation from occurring on the window

2. be sure to blank off the filter wheel and place the lens cap over the dewar window

POWER-DOWN SEQUENCE (lightning):

1. In the control room:

a. switch off the main power terminal and unplug it from AC power

b. in back of the PC, unplug the Ethernet and serial cables to “TCS instcon”

2. At the telescope,

a. switch off the Leach controller and unplug its power connector

USER INTERFACE:

a.	Type "i” to enter the miscellaneous menu:

1. type "i” to send "power on issued” command to controller

2. after ~10 seconds, "pon issued” appears in lower left of screen

3. type "d”;  to engage saving to SUN, type "m:” and return

to disable saving to BOK, enter a blank & return

4. type "b”;  to engage saving to PC, type the directory path

5. type "p”;  enter the <=5 character prefix name of your file

6. type "o”;  enter the name of your "offset” file

(EVEN if it is already printed)

7. type "q” and return to quit this menu

b. 	Type "d” to enter  the "display” menu:

1. type "r”  and "y” to start the display

2. type "q” and return to quit this menu

c.	Type "m” to enter the "movie” menu:

1. type "e”, set the  exposure time, and return

2. type "n”, set  the number of frames, and return

3. type "v” and "y” to  for save frames

[always use grab mode #3 & single reads]

4. type "q”  and return to quit this menu

d.	PISCES now takes an exposure in reset/read/integrate/read mode:

1. these messages should appear in the lower left screen:

MRA, binit, displaying, saving

2. DATA light on Leach interface box will blink for each read

3. a bell should sound when finished saving

e.	IF the PC hangs in MRA mode:

1. type "shift-Q” to un-hang the PC & display a garbage frame

2. press the "reset” button twice on the Leach interface box

3. take another frame

4. if the problem persists:

- power down the Leach interface & Leach controller

- repeat numbers 2,3,4 above

NEVER enter the "D” menu or attempt to change voltage settings on the detector.

OBSERVING:

DO NOT put bright sources on the detector for ANY length of time. These include the usual FK5 stars used at the 90-inch telescope for initial pointing calibration. The detector retains an imprint of bright objects for many hours or days.

DEWAR TILT: Done by Don to align pupil and cold stop in pupil-viewing mode

MMT ALIGNMENT:

1. establish rough collimation by positioning a star near center of array and

examining the out-of-focus image centration

2. measure center of rotation during a 180 sec exposure while rotator turns 180

deg. An 8th magnitude star works well. Rotator & PA angles = at 0 deg.

Use DS9 to fit concentric circles to star streaks.

10/13/02:	(745,526) on PC screen

12/19/02:	(751,514)

1/10/03:	(749,534)

3/10/03:	(758,512)

4/10/03:	(511,506)	using MMT Top Box mount

5/23/03:	(515,501)

12/6/03:	(510,514)

3. measure AZ/EL offsets between this center of rotation and acquisition location.

12/19/02: 	delta AZ =    0.3"; delta EL = 10.1"

1/10/03:	delta AZ =  12.4"; delta EL = 24.1"

3/10/03:	delta AZ =  46.8"; delta EL = -8.0"

5/23/03:	delta AZ = -21.8"; delta EL = 21.1"

12/6/03:	delta AZ =    5.4"; delta EL = -5.8"

4. measure RA/DEC offsets between center of rotation and desired centered

position

10/13/02:	delta RA = 0"; delta DEC = 40"

12/19/02:	delta RA = 0"; delta DEC = 37"

1/10/03:	delta RA = 4.1"; delta DEC = 43.8"

3/10/03:	delta RA = 16"; delta DEC = 48"

5. record hexapod values

12/19/02: 	hexapod values do not repeat after f/9 removal

6. engage open-loop correction of focus and coma vs. elevation angle

12/19/02: 	telescope operator said this was unreliable so we avoided it; had to refocus +40 z-units every +2 deg of elevation change

1/10/03:	“elcol(sp?)” engaged and works well

7. Orientation of N/E and AZ/EL inverts between northerly vs. southerly objects

Same occurs for direction of astigmatic images:

South objects: forward elongated (/) images, move hexapod down (smaller  numbers)

Calibrate: optical distortion by observing a dithered sequence on either an astrometric

field (e.g., globular cluster) or a crowded star field. The distortion map may

change between observing runs.

NOTES:

Exposure times have an 'overhead’ of 0.75 +/- 0.1  sec.

Non-linearity is 3% at 22000 ADU

SOFTWARE: The control program will take data whenever two conditions are both 'true’: display 'run (r)’ is on and the number (n) of frames in movie mode is positive.

SAVING: It is easy to overwrite data or to not record it at all. See comments below.

1. Type "h” to provide/edit header information:

a.	enter observer name, filter, site

b.	you have four lines of comment information available

NOTE: As long as the telescope control system (TCS) is operating, pointing information will be inserted into the header automatically. If you are taking dark frames after TCS is turned off, there may be residual garbage in the TCS data stream which may adversely affect the header.

2.  Type "m” to enter movie mode and set:

a.	your exposure time

b.	yes/no to saving

c.	if you do not wish to dither, enter the number of frames

d.	to dither a series of frames, do not set the number of frames, go 			to step 3 below

3.  To engage automatic dithering mode:

a. 	Disable "display” by typing "d” and "r” and answer no

b.	Type "i” to enter the miscellaneous menu:

1. be sure you have already entered your wobble file EVEN IF

the program already said 'offset.log’ at start-up

2. type "W” to engage wobbling:

message "wobbling engaged” appears in lower left

c.	Engage the display by typing "d” and "r” and answer yes

d. 	Off it goes!!

NOTES: 	The telescope will move after the image has been taken

There are only 18 allowable lines in the offset file.

You may want to adjust the “delay” parameter to allow for telescope settling

4. To abort while dithering:

a. 	Type a menu key like "m”

b.	Patiently wait for the current frame to finish

c.	Type "n” and zero the number of frames

d.	Start over

5.  Saving data:

a. 	Some users prefer to disable automatic saving to BOK. It is actually faster to ftp 			data from the PC to BOK outside of the "wfirc” program.

b. 	If you choose the ftp-method, you can re-enter "wfirc” and pick up the frame 			number where you left off:

1. type "i” to enter the miscellaneous menu

2. re-type the old prefix name

3. answer "yes” to ???? and enter the next number to append

c.	The PISCES program will NOT inform you if the PC drive is full or if you

are overwriting a file of the same name. Check the space available on the PC

hard drive every night and ftp as needed.

d.	If too many processes are running on the network, the saving process can hang.

6.  Other features relating to the TCS system:

a. you can "zero the differences” on the TCS screen to keep track of dithering or paddle 		motions of the telescope from the originally commanded position

b. you can wobble between two positions (beam #1 and beam #2) in RA and/or DEC

PC SCREEN:

Sky Orientation:

90":	North on the sky is to the right and East is down

(the screen orientation flips vertically on the SUN)

MMT:	with rotator angle at “position angle” =  zero degrees

northerly objects:

North on sky is right and East is down

+AZ on sky is up and +EL is right

southerly objects:

Opposite

in full frame mode, every other pixel is displayed

frame name refers to the NEXT frame to be saved

cursor box is 32x32 pixels but can be adjusted

cursor motions:

coarse jumps:	use keys 1,2,3,4

fine jumps: 	use arrow keys

The lower right panel plots a one-dimensional horizontal cut through the middle of the cursor display window immediately above. A software bug causes the  "retrace” of this cut to produce a black line diagonally across the screen to the left. This line is NOT in your data.

display shows:

the mean and sigma (in ADU) of pixels in the cursor box

the value (in ADU) of the central point of cursor box

the "cds” (correlated double sampling) value of central pixel

the "ped” (pedestal) value from the first read of central pixel

DETECTOR

[HAWAII HgCdTe with 1024x1024 pixels]

Pixel scale:

at 90”		0.4958 +/- 0.005 arcsec/pixel (Nalin Samarashinha measurement)

at MMT 	0.185 arcsec/pixel

Field diameter:

at 90”		8.5 arcmin

at MMT	3.1 arcmin

Gain:			~5.7 ADU/electron

Full well:		do not exceed ~20-25,000 ADU

NOTES:

Exposure times have an 'overhead’ of 0.75 +/- 0.1  sec.

Non-linearity is 3% at 22000 ADU.

Avoid putting bright sources on array for ANY length of time!

Bright sources (e.g. telescope FK5 stars):

can produce a "memory effect”

can produce a big "ped” value even though "cds” appears w/in full well

All sources leave "streaks” and "echoes” of cross-talk in all quadrants; this effect can be removed in data processing.

FILTERS

Be careful when changing filters:

The TWO gear systems must engage or else the filter number can change without the filter

wheel moving!!

A dark line, marked on the engager knob, should point downward to facilitate meshing the two

gear systems.

Be sure to retract the positioning rod to minimize thermal conduction.

Ten filter positions (Jan. 24, 2005):

1	2.035 um (WR-Cont1)

2	1203 nm	(need to check)	according to records from 11/9/04

3	1180 nm	(need to check)	according to records from 11/9/04

4	1145 nm	(need to check)	according to records from 11/9/04

5	open

6	J-continuum

7	H

8	Ks

9	J

10	BLANK (but it still leaks!!)

(When loading filters and looking down at box filter positions decrease CCW on the wheel.)

FOCUSING

The final f/-ratio on the detector is ~f/3.3, making focusing a critical adjustment.

The focus value depends on temperature filter position. Consult the focus database for estimates of expected focus and focus change with temperature.

Conduct a focus sequence on stars faint enough for a ~10 second exposure to avoid effects of seeing changes. At the 90”, I recommend an interval of 3 units between focus settings, always moving in the positive direction. At the MMT, 30 units.

If you have conducted a focusing sequence, please enter your new focus number, wavelength, and temperature (from the analog meter above the water cooler) into the log book.

The IRAF command "starfocus” (in the nmisc package) will examine multiple stars in your focus images, fit a function to the FWHM vs. focus, and display/print  your focus data and an interpolated best-focus value. At MMT, use a “fitting radius” of 15. You may want to change the "pixtype” to ushort first.

MMT (very) preliminary experience:

~5 units /degree of elevation; lower EL => smaller focus numbers

~20-40 units/degree of temperature; lower temp. => higher focus numbers

LIQUID NITROGEN

Your should replenish the liquid nitrogen (LN2) cryogen in the PISCES dewar at the beginning and end  of each night. Be sure to cover the dewar window while transferring cryogen.

Check that you have enough LN2 in the storage dewar. If the level is low, be sure to tell the day crew you need more LN2. They do not work on weekends so you have to think ahead.

There are two cryogen vessels. The outer (8 liters) shields the inner can (1 liter) which cools the detector and optics. The outer can has a higher evaporation rate because it is directly exposed to radiation for the ambient temperature case. A flexible hose is attached to each fill port to direct any overflows of LN2 away from the fragile dewar window.

The storage dewar is self-pressurized. Opening the main valve will cause LN2 to flow out the main pipe. If the storage dewar has recently been filled, you will need more pressure. In this case use the "bottle” of compressed nitrogen gas on the east wall to force nitrogen gas into the storage dewar at a low flow rate. You can couple the flexible  hose from the tank onto a side port of the storage dewar; then open the small valve on the storage dewar to begin the flow.

To fill the dewar:

1. SLIGHTLY LOOSEN the three LARGE bolts on each side of the purple dewar mount. The

dewar will tilt backwards.

2. Remove the flexible hoses over the fill ports

3. Using a funnel, pour LN2 into the two vessels. This may require as much as 8 liters.

4. Close all valves on storage dewar and compressed nitrogen gas tank.

5. Reattach the flexible hoses.

6. Tilt the dewar back up until the mount rests firmly against the micrometer head. Then tighten

all six bolts.

PREPARATION:

Prior to telescope run:

1. Print out PISCES log sheets from Joannah Hinz’s home page.

2. Make electronic offset files.

3. Make electronic object catalog.

The following are done during installation at the telescope. You should not have to repeat them.

1. PISCES computer:

a. 	set PC "date” & "time” with LOCAL values

b.	type "90inch.bat” to setup PCNFS paths/drives, then reboot

c.	"delete” & "rmdir” old data in directories on e: and f:

d.	"mkdir” new data directories on drives e: and f:

2. BOK SUN computer:

a.	"mkdir” the following data directory:

/d1/bokobs/rawdata/TEMP

b.	to display images on the separate monitor near BOK:

bring up an xterm window

setenv DISPLAY bok:0.1

ximtool & (this should be the only ximtool operating)

related commands to set screen dimensions:

set stdimage = imt1024

show stdimage

gdevices

set imtype = "fits”

may need to edit login.cl for imtype

3. Be sure detector (preamp & Leach controller), Leach interface box, & computer, are properly 	connected.

a. at 90”, you can use EITHER TCS1 or TCS2 connectors at patch panel

4. Measure readnoise frames from grounded inputs through:

1. Leach box only				(~1-1.3 ADU)

2. preamp + Leach box			(~1.3 ADU)

3. take a 30 sec dark frame for the 'library’	(~4 ADU at standard cursor position)

5. Record a frame for distortion library

MISCELLANEOUS

at 90": name of our PC on the KPNO network:	bokpcx (140.252.81.13)

at MMT: 	using PCNFS remotely mount a hard drive to 'ringo'

(login: mmtobs with password:  Mar5Attacks!)

name of additional xterm in control room:	ncdxtb22

Don McCarthy’s phone numbers:		742-7584 (H)

621-4079 (W) use "mailbox” #4

235-6604 (cell) I seldom use this

IMPROVEMENTS:

software:

LINUX system is the first priority for the future.

remove initial "offset.log” in menu

implement 'scripting’ for taking dark frames at end of night

variable number of co-adds during dithering

sense TCS statement of settling time after an offset/wobble

fix:	cursor retrace

make "D” mode off-limits

toggle quad display but save full frame

increase length of wobble files

new boot image for M51

detector voltage protection

hardware:

improve LN2 filling (lengthen the tubes to nearly touch bottom)

can we remove scattered light at K in "lens ring”?

multiple reads

subarrays

grism

CD writer

motorize filter wheel

loosen cover cap over entrance lenses

motorize pupil wheel

Software hints for Don:

mpage -f2Hm40 <file> > test.ps

lpr test.ps

unlearn disp

disp dev$pix

flpr

NOTES FOR DON:

In setting voltages, remember that the PC/Leach system takes 5-10 seconds to recover after quitting the shift-D menu. The system will blink the "data” light after that period so wait for the light before typing further instructions.

PUPIL CENTERING:

move micrometer to smaller numbers to move vertical spiders North (right) on PC screen

2/2/01:	micrometer = 030 with IIS PA = 90.9

MMT:

no nullmodem adaptor needed