## Slide 1
MAPS WFS Frontend Periscope Optical Design
Rev. 1.8
March 16, 2020

## Slide 2
Overview
Due to problems with procurement of reliable X-Y-Z stages for WFS patrol field acquisition and WFS nodding, alternative options were explored.
A Periscope, an OAP Relay, and an Offner Relay design were analyzed. The Periscope design with a linear Z-stage was the easiest and best option.
This slide set illustrates the worst case full patrol field cases and how a coordinated movement of five (5) DOF is required to boresight and focus the off-axis beam onto the WFS optical axis. View this presentation in ‘Slide Show’ mode.
The Periscope is a combination of the ARIES Dichroic and a Periscope Fold Mirror (PFM)
The PFM was sized based on the ACQ channel FOV of ± 10”

## Slide 3
Periscope Design
Top Box Floor
3” Fold Mirror
WFS Optical Axis
ARIES Dichroic
f/15 Focal Plane
± 10” FOV
There is not a lot of space to insert a periscope, so the ARIES Dichroic will be used as one of the fold mirrors.
The ARIES Dichroic is located ~183mm below the Top Box floor.
Placing the Periscope Fold Mirror 100mm downstream of the f/15 focal plane provides ~124mm of vertical clearance from the Top Box Floor for optical mounts and the breadboard.
~183 mm
~124 mm
MMT1.1 ARIES1.1 Stage2.5.zmx

## Slide 4
MMT1.1 ARIES1.1 Stage2.8a.zmx
3 DOF required
YZ Plane On-Axis Focus
Z-Piston
X-Tilt
X-Tilt
± 10mm
YZ Plane
XZ Plane

## Slide 5
MMT1.1 ARIES1.1 Stage2.8a.zmx
YZ Plane Boresight
2 DOF required
YZ Plane
XZ Plane
X-Tilt
X-Tilt
-1’
+1’

## Slide 6
MMT1.1 ARIES1.1 Stage2.11a.zmx
XZ Plane Boresight
2 DOF required
YZ Plane
XZ Plane
Y-Tilt
Y-Tilt
-1’
+1’

## Slide 7
YZ-Plane Boresight & Focus
MMT1.1 ARIES1.1 Stage2.8a.zmx
Z-Piston
X-Tilt
X-Tilt
3 DOF required
± 10mm
YZ Plane
XZ Plane

## Slide 8
XZ-Plane Boresight & Focus
MMT1.1 ARIES1.1 Stage2.11a.zmx
Z-Piston
X-Tilt
X-Tilt
5 DOF required
± 10mm
Y-Tilt
Y-Tilt
YZ Plane
XZ Plane

## Slide 9
Periscope DOF Specs
 | On-Axis | YZ Plane |  | XZ Plane |  | Specs
Alignment Type | Focus | Boresight | Boresight & Focus | Boresight | Boresight & Focus | -
DOF Required | 3 | 2 | 3 | 4 | 5 | 5
ARIES Dichroic X-Tilt [˚] | ± 1.4 | ± 2.4 | ± 3.9 | ± 0.2 | ± 1.5 | ± 3.9
ARIES Dichroic Y-Tilt [˚] | 0 | 0 | 0 | ± 2.4 | ± 2.5 | ± 2.5
Periscope Fold Mirror X-Tilt [˚] | ± 1.4 | ± 2.4 | ± 3.9 | ± 0.2 | ± 1.6 | ± 3.9
Periscope Fold Mirror Y-Tilt [˚] | 0 | 0 | 0 | ± 2.4 | ± 2.4 | ± 2.4
Periscope Fold Mirror Piston [mm] | ± 10 | 0 | ± 10 | 0 | ± 10 | ± 10

## Slide 10
3” Periscope Fold Mirror
ARIES Dichroic
f/15 Focal Plane
On-axis
MMT1.1 ARIES1.1 Stage2.5.zmx
+10” FOV
-10” FOV
ACQ Channel FOV

## Slide 11
Unfolded WFS Design
Base1.8 VPWFS4.7 IRPWFS1.3.2 ACQ2.4 Int5.3.zmx
Periscope
ACQ Channel
VPWFS Channel
IRPWFS Channel
Baseline
Telescope Dichroic
Periscope Fold Mirror
ITL
ADC
Modulator
BS
K-Mirror
Flip Mirror
CTL
CTL
Double Pyramid
Double Pyramid
f/15 FP
CCID75
Saphira
DCC3240N

## Slide 12
Top Box Shaded Model
Base1.8 VPWFS4.7 IRPWFS1.3.2 ACQ2.4 Int5.3.zmx

## Slide 13
Does It Fit?
Top Box Floor
ARIES Dichroic
Base1.8 VPWFS4.7 IRPWFS1.3.2 ACQ2.4 Int5.3.zmx
Periscope Fold Mirror
64” (1625.6 mm)
DCC3240N
CCID75
Saphira Cryostat